## Integration
Peak Integration Report
Sample Name: |  | 1 |  |  | Inj. Vol.: | 10
Injection Type: |  | Unknown |  |  | Dilution Factor: | 1
Instrument Method: |  | AS11HC_SM_UV |  |  | Operator: | labuser
Inj. Date / Time: |  | 2022-09-19 12:29 | 44823.5202 |  | Run Time: | 15.0033
No. | Time | Peak Name | Peak Type | Area | Height | Amount
| min |  |  | mAU*min | mAU | ppm
UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
1 | 4.3867 | Nitrite UV | BM | 0.20426 | 0.82592 | n.a.
4 | 6.3284 | Bromide UV | M | 0.042738 | 0.29935 | n.a.
5 | 6.6159 | Nitrate UV | MB | 140.7465 | 807.5825 | 42.0135
|  | TOTAL: |  | 140.9935 | 808.7078 | 42.0135

## Calibration
Calibration Batch Report
Sequence: | 091522 |  |  |  | Injection Volume: | 10
Instrument Method: | AS11HC_SM_UV |  |  |  | Operator: | labuser
Inj. Date / Time: | 2022-09-19 12:29 | 44823.5202 |  |  | Run Time: | 15.0033
Calibration Summary
Peak Name | Eval.Type | Cal.Type | Points | Offset | Slope | Curve | Coeff.Det.
|  |  |  | (C0) | (C1) | (C2) | %
UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
Nitrite UV | Area | Lin, WithOffset | 6 | 0.7459 | 3.2099 | 0 | 99.8632
Bromide UV | Area | Lin, WithOffset | 6 | 0.060011 | 0.32448 | 0 | 99.8724
Nitrate UV | Area | Lin, WithOffset | 6 | 0.7313 | 3.3326 | 0 | 99.9286
|  | AVERAGE: |  | 0.51241 | 2.289 | 0 | 99.8881
Injection Name | Ret.Time | Area | Height | Amount
| min | mAU*min | mAU | ppm
| UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
| Nitrite UV | Nitrite UV | Nitrite UV | Nitrite UV
1600x | 4.3985 | 0.22839 | 2.0269 | n.a.
800x | 4.4068 | 0.4559 | 4.0254 | n.a.
400x | 4.3985 | 0.92171 | 8.0363 | 0.054772
200x | 4.4027 | 1.8714 | 16.7136 | 0.35064
20x | 4.4069 | 18.2181 | 149.448 | 5.4432
5x | 4.4194 | 64.6011 | 487.5581 | 19.893
Average | 4.4055
Rel. Std. Dev. | 0.1762%
Injection Name | Ret.Time | Area | Height | Amount
| min | mAU*min | mAU | ppm
| UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
| Bromide UV | Bromide UV | Bromide UV | Bromide UV
1600x | 6.4277 | 0.025843 | 0.21862 | n.a.
800x | 6.4318 | 0.036713 | 0.34096 | n.a.
400x | 6.4318 | 0.071838 | 0.64622 | 0.036449
200x | 6.4319 | 0.14641 | 1.2651 | 0.26628
20x | 6.4152 | 1.8384 | 12.964 | 5.4806
5x | 6.3985 | 6.512 | 47.4131 | 19.8838
Average | 6.4228
Rel. Std. Dev. | 0.2107%
Injection Name | Ret.Time | Area | Height | Amount
| min | mAU*min | mAU | ppm
| UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
| Nitrate UV | Nitrate UV | Nitrate UV | Nitrate UV
1600x | 6.7068 | 0.45449 | 3.0428 | n.a.
800x | 6.7151 | 0.60324 | 3.9206 | n.a.
400x | 6.711 | 1.0103 | 6.4824 | 0.08372
200x | 6.711 | 2.0331 | 12.5803 | 0.39061
20x | 6.6902 | 17.9916 | 113.9728 | 5.1792
5x | 6.6569 | 67.2436 | 409.1535 | 19.9579
Average | 6.6985
Rel. Std. Dev. | 0.3312%

## Peak Analysis
Peak Analysis Report
Sample Name: |  | 1 |  |  |  | Injection Volume: |  | 10
Injection Type: |  | Unknown |  |  |  | Dilution Factor: |  | 1
Instrument Method: |  | AS11HC_SM_UV |  |  |  | Operator: |  | labuser
Inj. Date / Time: |  | 2022-09-19 12:29 | 44823.5202 |  |  | Run Time: |  | 15.0033
No. | Time | Peak Name | Width | Height | Resol. | Resol. | Plates | Plates | Asymmetry
| min |  | min | mAU | (USP) | (EP) | (USP) | (EP)
UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
1 | 4.3867 | Nitrite UV | 0.38534 | 0.82592 | 0.26758 | n.a. | 2074 | n.a. | n.a.
2 | 4.5867 |  | 1.1095 | 0.43289 | 0.58606 | n.a. | 273 | n.a. | n.a.
3 | 5.0325 |  | 0.41193 | 0.83751 | 0.45528 | n.a. | 2388 | 3439 | n.a.
4 | 6.3284 | Bromide UV | 5.2805 | 0.29935 | 0.10347 | n.a. | 23 | n.a. | n.a.
5 | 6.6159 | Nitrate UV | 0.27664 | 807.5825 | n.a. | n.a. | 9151 | 9399 | 1.3221
|  | AVERAGE: | 1.4928 | 161.9956 | 0.3531 | #DIV/0! | 2781.8 | 6419 | 1.3221

## Summary - INJ. vs ANION
Anion Summary Report
No. | Name | Dilution | Time | Area | Rel.Area | Height | Rel.Height | Amount | notes
|  |  | min | µS*min | % | µS | % | ppm
|  |  | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total
|  |  | Fluoride | Fluoride | Fluoride | Fluoride | Fluoride | Fluoride
1 | DI Water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 10 ppm Anion Std | 1 | 2.7633 | 3.0007 | 17.1096 | 33.3146 | 21.1451 | 5.029
3 | DI water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 1600x | 1 | 2.7633 | 0.015087 | 8.2384 | 0.10615 | 7.4655 | 0.0058817
5 | 800x | 1 | 2.7533 | 0.016628 | 6.3142 | 0.1881 | 8.677 | 0.008474
6 | 400x | 1 | 2.7633 | 0.029589 | 6.6255 | 0.37281 | 9.5371 | 0.030281
7 | 200x | 1 | 2.7567 | 0.063904 | 7.0945 | 0.82358 | 10.7056 | 0.088014
8 | 20x | 1 | 2.7633 | 0.6489 | 7.5849 | 7.9388 | 10.6919 | 1.0722
9 | 5x | 1 | 2.7633 | 2.3787 | 7.331 | 27.0009 | 10.1096 | 3.9826
10 | 10 ppm Anion Std | 1 | 2.7733 | 3.0003 | 17.0915 | 33.4246 | 21.1781 | 5.0283
11 | 062 | 1 | 2.7567 | 0.039479 | 0.24093 | 0.50565 | 0.37239 | 0.046921 | skip
12 | 055 | 1 | 2.76 | 0.093396 | 0.78009 | 1.1525 | 1.4393 | 0.13763 | skip
13 | 043 | 1 | 2.7633 | 0.024354 | 0.43414 | 0.30904 | 0.81213 | 0.021473 | skip
14 | 10 ppm Anion Std | 1 | 2.7633 | 2.9224 | 17.1966 | 32.1204 | 21.0462 | 4.8972
15 | 10 ppm Anion Std | 1 | 2.7667 | 2.9187 | 17.1757 | 32.1481 | 21.0602 | 4.8912
16 | 1 | 1 | 2.7603 | 0.030612 | 0.31828 | 0.39319 | 0.63143 | 0.032002
17 | 2 | 1 | 2.7567 | 0.033301 | 0.35647 | 0.41213 | 0.6795 | 0.036526
18 | 3 | 1 | 2.7633 | 0.03363 | 0.37081 | 0.33874 | 0.57273 | 0.037079
19 | 4 | 1 | 2.7467 | 0.034777 | 0.35585 | 0.44824 | 0.70918 | 0.039009
20 | 5 | 1 | 2.76 | 0.02878 | 0.29747 | 0.3563 | 0.56945 | 0.028921
21 | 6 | 1 | 2.76 | 0.027854 | 3.392 | 0.361 | 5.0221 | 0.027361
22 | 7 | 1 | 2.76 | 0.032603 | 4.0356 | 0.35272 | 5.0471 | 0.035352
23 | 8 | 1 | 2.7567 | 0.034198 | 4.2811 | 0.38867 | 5.6297 | 0.038036
24 | 9 | 1 | 2.7533 | 0.034232 | 4.0145 | 0.3933 | 5.3241 | 0.038093
25 | 10 | 1 | 2.7567 | 0.035708 | 4.2754 | 0.39255 | 5.4367 | 0.040576
26 | 11 | 1 | 2.76 | 0.040766 | 0.69457 | 0.51005 | 1.2264 | 0.049086
27 | 12 | 1 | 2.7567 | 0.041889 | 0.78031 | 0.55241 | 1.4317 | 0.050975
28 | 13 | 1 | 2.7567 | 0.048725 | 1.0425 | 0.60557 | 1.7957 | 0.062476
29 | 14 | 1 | 2.7567 | 0.069497 | 1.4906 | 0.87309 | 2.6106 | 0.097425
30 | 15 | 1 | 2.7533 | 0.053909 | 1.1743 | 0.68011 | 2.0385 | 0.071199
31 | 16 | 1 | 2.7633 | 0.038548 | 0.25765 | 0.46958 | 0.37394 | 0.045355
32 | 17 | 1 | 2.76 | 0.0405 | 0.25432 | 0.50449 | 0.37927 | 0.048638
33 | 18 | 1 | 2.76 | 0.027877 | 0.19835 | 0.36508 | 0.3085 | 0.0274
34 | 19 | 1 | 2.7633 | 0.036626 | 0.28516 | 0.44162 | 0.40621 | 0.042121
35 | 20 | 1 | 2.7567 | 0.042551 | 0.31913 | 0.5102 | 0.45314 | 0.052089
36 | 21 | 1 | 2.81 | 0.0054288 | 5.7016 | 0.067188 | 10.9915 | n.a.
37 | A-5x | 1 | 2.7633 | 2.374 | 7.3264 | 26.6805 | 10.0156 | 3.9747
38 | MQ Blank 1 | 1 | 2.8067 | 0.0053747 | 6.0398 | 0.051277 | 11.0458 | n.a.
39 | 22 | 1 | 2.7667 | 0.041089 | 0.39381 | 0.46794 | 0.69865 | 0.049629
40 | 23 | 1 | 2.7533 | 0.038812 | 0.36529 | 0.46597 | 0.69002 | 0.045798
41 | 24 | 1 | 2.7567 | 0.041756 | 0.39127 | 0.44426 | 0.65631 | 0.050752
42 | 25 | 1 | 2.7533 | 0.037346 | 0.35514 | 0.47287 | 0.70515 | 0.043331
43 | 26 | 1 | 2.7533 | 0.038467 | 0.36823 | 0.46268 | 0.69597 | 0.045218
44 | 27 | 1 | 2.7533 | 0.024592 | 2.9633 | 0.2978 | 4.1254 | 0.021875
45 | 28 | 1 | 2.7667 | 0.023931 | 2.8639 | 0.28701 | 3.9577 | 0.020761
46 | 29 | 1 | 2.7567 | 0.027391 | 3.2652 | 0.32308 | 4.4666 | 0.026582
47 | 30 | 1 | 2.76 | 0.026951 | 3.0055 | 0.31896 | 4.0957 | 0.025843
48 | 31 | 1 | 2.7567 | 0.026133 | 3.0403 | 0.34262 | 4.5934 | 0.024467
49 | 32 | 1 | 2.7633 | 0.08841 | 1.1269 | 1.1427 | 2.1081 | 0.12924
50 | 33 | 1 | 2.76 | 0.088631 | 1.2724 | 1.1096 | 2.2801 | 0.12962
51 | 34 | 1 | 2.7567 | 0.104 | 1.5184 | 1.1631 | 2.432 | 0.15547
52 | 35 | 1 | 2.7567 | 0.098273 | 1.7027 | 1.1835 | 3.055 | 0.14584
53 | 36 | 1 | 2.7567 | 0.090834 | 1.5172 | 1.1749 | 2.8418 | 0.13332
54 | 37 | 1 | 2.76 | 0.053384 | 0.28577 | 0.67968 | 0.44348 | 0.070314
55 | 5 ppm F/10 ppm Anion Std | 1 | 2.7667 | 2.9241 | 17.1726 | 32.3656 | 21.1025 | 4.9002
56 | 38 | 1 | 2.7533 | 0.056421 | 0.31058 | 0.67575 | 0.45035 | 0.075425
57 | 39 | 1 | 2.76 | 0.053574 | 0.28586 | 0.63385 | 0.41205 | 0.070634
58 | 40 | 1 | 2.76 | 0.053382 | 0.27469 | 0.69307 | 0.43616 | 0.070312
59 | 41 | 1 | 2.7533 | 0.062967 | 0.29573 | 0.81032 | 0.47054 | 0.086438
60 | 42 | 1 | 2.7867 | 0.017631 | 13.4678 | 0.19292 | 20.2723 | 0.010162
61 | A-400x | 1 | 2.76 | 0.027384 | 6.2802 | 0.36152 | 9.4881 | 0.026572
62 | MQ Blank 2 | 1 | 2.8 | 0.0049082 | 5.7286 | 0.048398 | 10.1765 | n.a.
63 | 43 | 1 | 2.7533 | 0.040313 | 0.37352 | 0.48434 | 0.70976 | 0.048324
64 | 44 | 1 | 2.7533 | 0.034682 | 0.31946 | 0.44144 | 0.64227 | 0.03885
65 | 45 | 1 | 2.76 | 0.037771 | 0.37624 | 0.4781 | 0.7422 | 0.044047
66 | 46 | 1 | 2.76 | 0.045087 | 0.70698 | 0.48308 | 1.1245 | 0.056355
67 | 47 | 1 | 2.7533 | 0.037915 | 0.51386 | 0.41604 | 0.84926 | 0.044289
68 | 48 | 1 | 2.7567 | 0.02587 | 2.8103 | 0.33412 | 4.1615 | 0.024023
69 | 49 | 1 | 2.76 | 0.023272 | 2.7653 | 0.3209 | 4.3155 | 0.019653
70 | 50 | 1 | 2.7533 | 0.026284 | 3.1434 | 0.30217 | 4.1177 | 0.02472
71 | 51 | 1 | 2.7533 | 0.026157 | 3.0411 | 0.30635 | 4.0646 | 0.024507
72 | 52 | 1 | 2.7533 | 0.024427 | 2.8745 | 0.31926 | 4.3069 | 0.021596
73 | 53 | 1 | 2.7533 | 0.098714 | 1.8471 | 1.2291 | 3.2188 | 0.14658
74 | 54 | 1 | 2.7567 | 0.099699 | 2.6078 | 1.2897 | 4.5829 | 0.14824
75 | 55 | 1 | 2.76 | 0.082097 | 1.1439 | 1.0382 | 2.0757 | 0.11862
76 | 56 | 1 | 2.76 | 0.10418 | 2.2137 | 1.2562 | 3.6818 | 0.15577
77 | 57 | 1 | 2.7633 | 0.085999 | 2.3115 | 1.0697 | 4.0953 | 0.12519
78 | 58 | 1 | 2.7533 | 0.046514 | 0.20826 | 0.60131 | 0.33455 | 0.058756
79 | 59 | 1 | 2.76 | 0.048944 | 0.22169 | 0.60565 | 0.34083 | 0.062845
80 | 60 | 1 | 2.7533 | 0.048722 | 0.25131 | 0.59664 | 0.37712 | 0.062471
81 | 61 | 1 | 2.7533 | 0.053248 | 0.25086 | 0.63901 | 0.37146 | 0.070087
82 | 62 | 1 | 2.76 | 0.052248 | 0.26076 | 0.63192 | 0.38699 | 0.068404
83 | 63 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
84 | A-200x | 1 | 2.7533 | 0.061731 | 6.9611 | 0.80924 | 10.7508 | 0.084359
85 | MQ Blank 3 | 1 | 2.8 | 0.005421 | 4.8426 | 0.052199 | 10.4893 | n.a.
86 | 64 | 1 | 2.7667 | 0.056083 | 0.95058 | 0.57623 | 1.504 | 0.074856
87 | 65 | 1 | 2.7667 | 0.055458 | 0.89604 | 0.58439 | 1.4446 | 0.073805
88 | 66 | 1 | 2.7533 | 0.040024 | 0.66116 | 0.48011 | 1.2094 | 0.047837
89 | 67 | 1 | 2.7567 | 0.05039 | 0.83294 | 0.52428 | 1.3308 | 0.065278
90 | 68 | 1 | 2.76 | 0.049572 | 0.79978 | 0.50906 | 1.2365 | 0.063902
91 | 69 | 1 | 2.7767 | 0.080885 | 7.2755 | 0.80313 | 8.3561 | 0.11658
92 | 70 | 1 | 2.78 | 0.085236 | 7.6192 | 0.83932 | 8.728 | 0.1239
93 | 71 | 1 | 2.7833 | 0.092204 | 8.0389 | 0.87565 | 8.9024 | 0.13563
94 | 72 | 1 | 2.7767 | 0.067762 | 6.3273 | 0.66143 | 7.2308 | 0.094505
95 | 73 | 1 | 2.7833 | 0.087837 | 7.9277 | 0.85917 | 8.9545 | 0.12828
96 | 74 | 1 | 2.7667 | 0.098867 | 2.1581 | 1.0425 | 3.4694 | 0.14684
97 | 75 | 1 | 2.8133 | 0.49216 | 9.3532 | 3.1384 | 9.3058 | 0.80853
98 | 76 | 1 | 2.7567 | 0.11439 | 2.1425 | 1.0772 | 2.9946 | 0.17296
99 | 77 | 1 | 2.7567 | 0.090358 | 2.0915 | 0.94835 | 3.43 | 0.13252
100 | 78 | 1 | 2.76 | 0.10499 | 1.9301 | 1.0888 | 2.966 | 0.15713
101 | 79 | 1 | 2.7533 | 0.064607 | 0.37746 | 0.7707 | 0.54789 | 0.089198
102 | 80 | 1 | 2.7533 | 0.054426 | 0.28853 | 0.69828 | 0.45445 | 0.072068
103 | 81 | 1 | 2.7533 | 0.058502 | 0.27369 | 0.75302 | 0.4376 | 0.078926
104 | 82 | 1 | 2.7567 | 0.061138 | 0.33465 | 0.62934 | 0.42081 | 0.08336
105 | 83 | 1 | 2.7533 | 0.043744 | 0.26864 | 0.56146 | 0.41838 | 0.054096
106 | 84 | 1 | 2.8033 | 0.0066975 | 7.7545 | 0.076334 | 14.094 | n.a.
107 | A-800x | 1 | 2.7567 | 0.012998 | 5.2608 | 0.17321 | 8.332 | 0.0023668
108 | MQ Blank 4 | 1 | 2.79 | 0.0053054 | 5.2902 | 0.051019 | 10.5979 | n.a.
109 | SB 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
110 | SB 2 | 1 | 2.8 | 0.003817 | 4.7674 | 0.045144 | 9.8322 | n.a.
111 | SB 3 | 1 | 2.7667 | 0.0002792 | 0.37074 | 0.0035333 | 0.89015 | n.a.
112 | SB 4 | 1 | 2.7767 | 0.0002886 | 0.38969 | 0.0033553 | 0.86109 | n.a.
113 | 5 ppm F/10 ppm Anion Std | 1 | 2.7633 | 2.9198 | 17.1558 | 32.4609 | 21.156 | 4.893
| Sum: |  | 301.187 | 28.2611 | 370.3126 | 315.6102 | 522.4384 | 45.5347
| Average: |  | 2.7632 | 0.25928 | 3.3974 | 2.8955 | 4.793 | 0.45535
| Rel.Std.Dev: |  | 0.4832% | 280.4% | 126.9% | 278.7% | 115.7% | 279%
No. | Name | Dilution | Time | Area | Rel.Area | Height | Rel.Height | Amount
|  |  | min | µS*min | % | µS | % | ppm
|  |  | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total
|  |  | Chloride | Chloride | Chloride | Chloride | Chloride | Chloride
1 | DI Water | 1 | 3.7968 | 0.00072747 | 2.0976 | 0.0071516 | 3.2491 | n.a.
2 | 10 ppm Anion Std | 1 | 3.8167 | 3.7954 | 21.6412 | 41.8996 | 26.5941 | 10.3709
3 | DI water | 1 | 3.8 | 0.00067465 | 1.8813 | 0.0076162 | 3.3971 | n.a.
4 | 1600x | 1 | 3.8167 | 0.025192 | 13.7564 | 0.27804 | 19.5552 | n.a.
5 | 800x | 1 | 3.8133 | 0.043295 | 16.4409 | 0.48702 | 22.4664 | n.a.
6 | 400x | 1 | 3.82 | 0.086701 | 19.4141 | 0.9937 | 25.4209 | 0.1139
7 | 200x | 1 | 3.8133 | 0.18727 | 20.79 | 2.1233 | 27.6011 | 0.39203
8 | 20x | 1 | 3.82 | 1.8748 | 21.9141 | 20.8181 | 28.0376 | 5.0591
9 | 5x | 1 | 3.82 | 7.2725 | 22.4131 | 79.6278 | 29.8138 | 19.9873
10 | 10 ppm Anion Std | 1 | 3.8233 | 3.7965 | 21.6274 | 41.9187 | 26.5602 | 10.374
11 | 062 | 1 | 3.8167 | 0.69289 | 4.2285 | 7.7233 | 5.688 | 1.7904 | skip
12 | 055 | 1 | 3.82 | 1.0297 | 8.6008 | 11.3853 | 14.2185 | 2.722 | skip
13 | 043 | 1 | 3.82 | 0.088182 | 1.572 | 1.0082 | 2.6494 | 0.118 | skip
14 | 10 ppm Anion Std | 1 | 3.8167 | 3.689 | 21.7079 | 40.7268 | 26.6854 | 10.0766
15 | 10 ppm Anion Std | 1 | 3.8167 | 3.6871 | 21.6971 | 40.6853 | 26.6528 | 10.0713
16 | 1 | 1 | 3.817 | 0.14896 | 1.5488 | 1.6953 | 2.7225 | 0.28609
17 | 2 | 1 | 3.8133 | 0.14916 | 1.5967 | 1.6967 | 2.7974 | 0.28664
18 | 3 | 1 | 3.8167 | 0.13892 | 1.5318 | 1.5935 | 2.6942 | 0.25833
19 | 4 | 1 | 3.8 | 0.15589 | 1.5951 | 1.7714 | 2.8026 | 0.30525
20 | 5 | 1 | 3.82 | 0.14944 | 1.5445 | 1.6981 | 2.7139 | 0.28741
21 | 6 | 1 | 3.8167 | 0.10808 | 13.1616 | 1.2355 | 17.1871 | 0.17303
22 | 7 | 1 | 3.8167 | 0.095868 | 11.8666 | 1.0994 | 15.7313 | 0.13926
23 | 8 | 1 | 3.81 | 0.096969 | 12.1389 | 1.111 | 16.0923 | 0.1423
24 | 9 | 1 | 3.81 | 0.10562 | 12.3867 | 1.2094 | 16.3714 | 0.16624
25 | 10 | 1 | 3.81 | 0.10848 | 12.988 | 1.2382 | 17.1485 | 0.17412
26 | 11 | 1 | 3.8167 | 0.33275 | 5.6694 | 3.7414 | 8.996 | 0.7944
27 | 12 | 1 | 3.8133 | 0.35251 | 6.5666 | 3.9683 | 10.285 | 0.84903
28 | 13 | 1 | 3.81 | 0.33071 | 7.0759 | 3.7162 | 11.0195 | 0.78874
29 | 14 | 1 | 3.81 | 0.35705 | 7.6584 | 4.0067 | 11.9804 | 0.8616
30 | 15 | 1 | 3.81 | 0.37788 | 8.2314 | 4.2374 | 12.7009 | 0.9192
31 | 16 | 1 | 3.8233 | 0.71075 | 4.7506 | 7.9383 | 6.3214 | 1.8398
32 | 17 | 1 | 3.82 | 0.70562 | 4.431 | 7.8856 | 5.9283 | 1.8256
33 | 18 | 1 | 3.8167 | 0.66159 | 4.7074 | 7.4073 | 6.2595 | 1.7038
34 | 19 | 1 | 3.82 | 0.62605 | 4.8742 | 6.9934 | 6.4326 | 1.6056
35 | 20 | 1 | 3.8133 | 0.66221 | 4.9665 | 7.3975 | 6.5703 | 1.7056
36 | 21 | 1 | 3.8167 | 0.0063416 | 6.6602 | 0.072085 | 11.7926 | n.a.
37 | A-5x | 1 | 3.8233 | 7.2798 | 22.4659 | 79.6609 | 29.9038 | 20.0073
38 | MQ Blank 1 | 1 | 3.81 | 0.0031713 | 3.5637 | 0.020045 | 4.3181 | n.a.
39 | 22 | 1 | 3.82 | 0.16008 | 1.5342 | 1.8382 | 2.7445 | 0.31685
40 | 23 | 1 | 3.81 | 0.15838 | 1.4906 | 1.7959 | 2.6594 | 0.31214
41 | 24 | 1 | 3.8133 | 0.1624 | 1.5218 | 1.8373 | 2.7144 | 0.32327
42 | 25 | 1 | 3.81 | 0.16213 | 1.5418 | 1.8394 | 2.7429 | 0.32252
43 | 26 | 1 | 3.81 | 0.15955 | 1.5273 | 1.8091 | 2.7213 | 0.31538
44 | 27 | 1 | 3.81 | 0.098912 | 11.9185 | 1.1316 | 15.6758 | 0.14767
45 | 28 | 1 | 3.8233 | 0.10092 | 12.0778 | 1.1549 | 15.9254 | 0.15323
46 | 29 | 1 | 3.81 | 0.098922 | 11.7924 | 1.1325 | 15.6573 | 0.1477
47 | 30 | 1 | 3.8133 | 0.10286 | 11.4708 | 1.1775 | 15.1204 | 0.1586
48 | 31 | 1 | 3.8133 | 0.09894 | 11.5105 | 1.1306 | 15.158 | 0.14775
49 | 32 | 1 | 3.82 | 0.53562 | 6.827 | 5.9725 | 11.0177 | 1.3555
50 | 33 | 1 | 3.8167 | 0.48906 | 7.0212 | 5.4627 | 11.2251 | 1.2267
51 | 34 | 1 | 3.81 | 0.4908 | 7.1659 | 5.4689 | 11.4354 | 1.2315
52 | 35 | 1 | 3.8133 | 0.42084 | 7.2917 | 4.6803 | 12.0818 | 1.038
53 | 36 | 1 | 3.8133 | 0.43037 | 7.1883 | 4.8603 | 11.756 | 1.0644
54 | 37 | 1 | 3.8167 | 0.73832 | 3.9523 | 8.2373 | 5.3747 | 1.916
55 | 5 ppm F/10 ppm Anion Std | 1 | 3.8167 | 3.6867 | 21.6507 | 40.8637 | 26.6433 | 10.0701
56 | 38 | 1 | 3.81 | 0.72443 | 3.9878 | 8.1116 | 5.4059 | 1.8776
57 | 39 | 1 | 3.8133 | 0.72273 | 3.8564 | 8.0777 | 5.2512 | 1.8729
58 | 40 | 1 | 3.8167 | 0.73149 | 3.7641 | 8.1775 | 5.1462 | 1.8972
59 | 41 | 1 | 3.81 | 0.76902 | 3.6118 | 8.5822 | 4.9836 | 2.0009
60 | 42 | 1 | 3.81 | 0.013934 | 10.6438 | 0.15979 | 16.7905 | n.a.
61 | A-400x | 1 | 3.8167 | 0.085378 | 19.5803 | 0.98311 | 25.8021 | 0.11025
62 | MQ Blank 2 | 1 | 3.8067 | 0.0016876 | 1.9697 | 0.018426 | 3.8744 | n.a.
63 | 43 | 1 | 3.81 | 0.16996 | 1.5748 | 1.9302 | 2.8286 | 0.34418
64 | 44 | 1 | 3.81 | 0.17053 | 1.5707 | 1.936 | 2.8167 | 0.34573
65 | 45 | 1 | 3.8167 | 0.17627 | 1.7559 | 1.995 | 3.097 | 0.36162
66 | 46 | 1 | 3.8133 | 0.1059 | 1.6606 | 1.2041 | 2.8028 | 0.16701
67 | 47 | 1 | 3.8067 | 0.12061 | 1.6346 | 1.3744 | 2.8056 | 0.20769
68 | 48 | 1 | 3.8133 | 0.11519 | 12.5134 | 1.3168 | 16.4007 | 0.19269
69 | 49 | 1 | 3.8167 | 0.10104 | 12.0059 | 1.1626 | 15.6347 | 0.15356
70 | 50 | 1 | 3.81 | 0.10623 | 12.7053 | 1.2184 | 16.6035 | 0.16793
71 | 51 | 1 | 3.81 | 0.10361 | 12.046 | 1.1892 | 15.7783 | 0.16067
72 | 52 | 1 | 3.81 | 0.099804 | 11.7445 | 1.1459 | 15.4582 | 0.15014
73 | 53 | 1 | 3.81 | 0.61118 | 11.4363 | 6.813 | 17.8423 | 1.5644
74 | 54 | 1 | 3.8167 | 0.6189 | 16.1886 | 6.8929 | 24.4943 | 1.5858
75 | 55 | 1 | 3.8167 | 0.59124 | 8.2383 | 6.6059 | 13.208 | 1.5093
76 | 56 | 1 | 3.82 | 0.54034 | 11.482 | 6.0277 | 17.6673 | 1.3685
77 | 57 | 1 | 3.82 | 0.30439 | 8.1815 | 3.4274 | 13.1213 | 0.71594
78 | 58 | 1 | 3.81 | 0.78276 | 3.5046 | 8.7501 | 4.8682 | 2.0389
79 | 59 | 1 | 3.8167 | 0.76135 | 3.4485 | 8.5091 | 4.7884 | 1.9797
80 | 60 | 1 | 3.8133 | 0.63513 | 3.2761 | 7.105 | 4.4909 | 1.6307
81 | 61 | 1 | 3.81 | 0.7456 | 3.5126 | 8.339 | 4.8475 | 1.9362
82 | 62 | 1 | 3.8167 | 0.72752 | 3.6309 | 8.1316 | 4.9799 | 1.8862
83 | 63 | 1 | 3.8133 | 0.0031011 | 3.2497 | 0.035532 | 6.3824 | n.a.
84 | A-200x | 1 | 3.8067 | 0.18636 | 21.0145 | 2.1186 | 28.1461 | 0.38952
85 | MQ Blank 3 | 1 | 3.8033 | 0.0042125 | 3.7631 | 0.0213 | 4.2802 | n.a.
86 | 64 | 1 | 3.8167 | 0.27076 | 4.5893 | 3.0247 | 7.8948 | 0.62295
87 | 65 | 1 | 3.8167 | 0.27008 | 4.3637 | 3.0199 | 7.4653 | 0.62106
88 | 66 | 1 | 3.81 | 0.25024 | 4.1339 | 2.8108 | 7.0807 | 0.56621
89 | 67 | 1 | 3.81 | 0.25625 | 4.2358 | 2.8631 | 7.2677 | 0.58282
90 | 68 | 1 | 3.8133 | 0.25258 | 4.075 | 2.8241 | 6.8596 | 0.57266
91 | 69 | 1 | 3.81 | 0.1999 | 17.9806 | 2.2654 | 23.5696 | 0.42697
92 | 70 | 1 | 3.8133 | 0.20169 | 18.0288 | 2.2864 | 23.7759 | 0.43192
93 | 71 | 1 | 3.8133 | 0.20276 | 17.6777 | 2.2942 | 23.3244 | 0.43488
94 | 72 | 1 | 3.8133 | 0.1706 | 15.9294 | 1.939 | 21.1969 | 0.34592
95 | 73 | 1 | 3.8133 | 0.19441 | 17.5467 | 2.2045 | 22.9763 | 0.41179
96 | 74 | 1 | 3.82 | 0.59209 | 12.9241 | 6.5522 | 21.8054 | 1.5116
97 | 75 | 1 | 3.8133 | 0.63403 | 12.0495 | 6.9983 | 20.7509 | 1.6276
98 | 76 | 1 | 3.8133 | 0.66036 | 12.3679 | 7.3064 | 20.3123 | 1.7004
99 | 77 | 1 | 3.8133 | 0.47978 | 11.1057 | 5.3292 | 19.2748 | 1.201
100 | 78 | 1 | 3.8133 | 0.76143 | 13.9986 | 8.4148 | 22.9229 | 1.98
101 | 79 | 1 | 3.81 | 0.72491 | 4.2352 | 8.0834 | 5.7464 | 1.879
102 | 80 | 1 | 3.81 | 0.73878 | 3.9165 | 8.2371 | 5.3608 | 1.9173
103 | 81 | 1 | 3.81 | 0.79977 | 3.7415 | 8.8905 | 5.1665 | 2.086
104 | 82 | 1 | 3.81 | 0.72476 | 3.9671 | 8.0995 | 5.4157 | 1.8785
105 | 83 | 1 | 3.81 | 0.60361 | 3.7068 | 6.7481 | 5.0285 | 1.5435
106 | 84 | 1 | 3.81 | 0.0039495 | 4.5728 | 0.046293 | 8.5474 | n.a.
107 | A-800x | 1 | 3.8133 | 0.041073 | 16.6242 | 0.47606 | 22.9005 | n.a.
108 | MQ Blank 4 | 1 | 3.7933 | 0.003392 | 3.3823 | 0.014604 | 3.0336 | n.a.
109 | SB 1 | 1 | 3.8067 | 0.021049 | 17.3577 | 0.24414 | 30.1656 | n.a.
110 | SB 2 | 1 | 3.8 | 0.00093125 | 1.1631 | 0.010778 | 2.3474 | n.a.
111 | SB 3 | 1 | 3.81 | 0.0011037 | 1.4656 | 0.01295 | 3.2626 | n.a.
112 | SB 4 | 1 | 3.81 | 0.00086307 | 1.1654 | 0.010137 | 2.6015 | n.a.
113 | 5 ppm F/10 ppm Anion Std | 1 | 3.8133 | 3.6814 | 21.6307 | 40.8142 | 26.6002 | 10.0556
| Sum: |  | 430.8872 | 71.6011 | 984.7331 | 794.6572 | 1391.2639 | 185.4551
| Average: |  | 3.8132 | 0.63364 | 8.7145 | 7.0324 | 12.3121 | 1.9318
| Rel.Std.Dev: |  | 0.1386% | 191.8% | 75.69% | 189.8% | 69.9% | 184.4%
No. | Name | Dilution | Time | Area | Rel.Area | Height | Rel.Height | Amount
|  |  | min | mAU*min | % | mAU | % | ppm
|  |  | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
|  |  | Nitrite UV | Nitrite UV | Nitrite UV | Nitrite UV | Nitrite UV | Nitrite UV
1 | DI Water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 10 ppm Anion Std | 1 | 4.4151 | 34.9529 | 45.2463 | 277.5321 | 52.1921 | 10.6566
3 | DI water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 1600x | 1 | 4.3985 | 0.22839 | 29.3176 | 2.0269 | 35.618 | n.a.
5 | 800x | 1 | 4.4068 | 0.4559 | 37.2621 | 4.0254 | 45.325 | n.a.
6 | 400x | 1 | 4.3985 | 0.92171 | 38.4616 | 8.0363 | 47.3823 | 0.054772
7 | 200x | 1 | 4.4027 | 1.8714 | 37.7475 | 16.7136 | 49.4133 | 0.35064
8 | 20x | 1 | 4.4069 | 18.2181 | 44.4174 | 149.448 | 52.8585 | 5.4432
9 | 5x | 1 | 4.4194 | 64.6011 | 44.8518 | 487.5581 | 50.9395 | 19.893
10 | 10 ppm Anion Std | 1 | 4.4069 | 34.9558 | 45.2946 | 278.1019 | 52.2425 | 10.6575
11 | 062 | 1 | 4.386 | 0.2234 | 26.0595 | 0.80123 | 38.9864 | n.a. | skip
12 | 055 | 1 | 4.4152 | 0.93625 | 0.67099 | 7.5174 | 0.93672 | 0.059299 | skip
13 | 043 | 1 | 4.3985 | 0.12278 | 0.1542 | 1.223 | 0.25101 | n.a. | skip
14 | 10 ppm Anion Std | 1 | 4.411 | 34.2467 | 44.8731 | 271.1422 | 51.4458 | 10.4366
15 | 10 ppm Anion Std | 1 | 4.4193 | 34.1981 | 44.9055 | 271.1308 | 51.4625 | 10.4215
16 | 1 | 1 | 4.3867 | 0.20426 | 0.14424 | 0.82592 | 0.10197 | n.a.
17 | 2 | 1 | 4.3992 | 0.22705 | 0.16537 | 0.91055 | 0.11547 | n.a.
18 | 3 | 1 | 4.3742 | 0.20521 | 0.15278 | 0.78299 | 0.10117 | n.a.
19 | 4 | 1 | 4.3867 | 0.24697 | 0.17199 | 1.0389 | 0.1267 | n.a.
20 | 5 | 1 | 4.3909 | 0.22635 | 0.15872 | 0.88212 | 0.10819 | n.a.
21 | 6 | 1 | 4.3075 | 0.15974 | 4.7482 | 0.60905 | 3.1319 | n.a.
22 | 7 | 1 | 4.32 | 0.15909 | 4.3636 | 0.5958 | 2.8245 | n.a.
23 | 8 | 1 | 4.3033 | 0.15239 | 3.7217 | 0.60078 | 2.7543 | n.a.
24 | 9 | 1 | 4.3075 | 0.16732 | 4.2901 | 0.59511 | 2.621 | n.a.
25 | 10 | 1 | 4.3158 | 0.15713 | 4.2298 | 0.60696 | 3.0134 | n.a.
26 | 11 | 1 | 4.395 | 0.19251 | 0.26807 | 0.83513 | 0.18931 | n.a.
27 | 12 | 1 | 4.395 | 0.21448 | 0.33191 | 0.94176 | 0.23366 | n.a.
28 | 13 | 1 | 4.4075 | 0.22423 | 0.41155 | 0.89433 | 0.26274 | n.a.
29 | 14 | 1 | 4.4075 | 0.23307 | 0.46329 | 0.93788 | 0.29562 | n.a.
30 | 15 | 1 | 4.4033 | 0.24384 | 0.48742 | 1.1648 | 0.37053 | n.a.
31 | 16 | 1 | 4.3867 | 0.20092 | 13.9384 | 0.71447 | 14.6696 | n.a.
32 | 17 | 1 | 4.3866 | 0.20487 | 11.4788 | 0.83928 | 14.9115 | n.a.
33 | 18 | 1 | 4.3492 | 0.17951 | 10.9583 | 0.67441 | 11.7413 | n.a.
34 | 19 | 1 | 4.3325 | 0.1549 | 5.3506 | 0.61676 | 4.7938 | n.a.
35 | 20 | 1 | 4.3825 | 0.17519 | 7.2176 | 0.67626 | 7.0384 | n.a.
36 | 21 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
37 | A-5x | 1 | 4.4159 | 65.0514 | 44.4174 | 490.125 | 50.6366 | 20.0333
38 | MQ Blank 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
39 | 22 | 1 | 4.3908 | 0.50049 | 0.33315 | 2.7634 | 0.32268 | n.a.
40 | 23 | 1 | 4.4034 | 0.27668 | 0.18091 | 1.2814 | 0.14803 | n.a.
41 | 24 | 1 | 4.4033 | 0.55672 | 0.36282 | 3.177 | 0.36552 | n.a.
42 | 25 | 1 | 4.3991 | 0.36368 | 0.23966 | 2.0704 | 0.24087 | n.a.
43 | 26 | 1 | 4.3992 | 0.47811 | 0.3181 | 2.5973 | 0.30429 | n.a.
44 | 27 | 1 | 4.3158 | 0.15065 | 4.6138 | 0.60479 | 3.228 | n.a.
45 | 28 | 1 | 4.3117 | 0.16768 | 5.2157 | 0.61458 | 3.3469 | n.a.
46 | 29 | 1 | 4.3117 | 0.15032 | 4.3365 | 0.61003 | 3.0082 | n.a.
47 | 30 | 1 | 4.3116 | 0.16975 | 4.0997 | 0.6184 | 2.7861 | n.a.
48 | 31 | 1 | 4.3159 | 0.17028 | 4.2127 | 0.62244 | 3.0622 | n.a.
49 | 32 | 1 | 4.4075 | 0.32687 | 0.34736 | 1.6796 | 0.2974 | n.a.
50 | 33 | 1 | 4.4034 | 0.38456 | 0.48102 | 2.2741 | 0.46353 | n.a.
51 | 34 | 1 | 4.4075 | 1.2517 | 1.6312 | 9.7499 | 2.0517 | 0.15757
52 | 35 | 1 | 4.4242 | 0.74168 | 1.3985 | 5.9964 | 1.779 | n.a.
53 | 36 | 1 | 4.4159 | 1.0644 | 1.8137 | 8.3006 | 2.2392 | 0.099229
54 | 37 | 1 | 4.3658 | 0.17975 | 4.4317 | 0.69448 | 3.4649 | n.a.
55 | 5 ppm F/10 ppm Anion Std | 1 | 4.4117 | 34.2346 | 45.113 | 273.1993 | 51.6847 | 10.4328
56 | 38 | 1 | 4.3659 | 0.16205 | 4.5387 | 0.61755 | 3.5064 | n.a.
57 | 39 | 1 | 4.3867 | 0.21954 | 12.129 | 0.97112 | 13.2444 | n.a.
58 | 40 | 1 | 4.395 | 0.28144 | 14.961 | 1.4797 | 19.3375 | n.a.
59 | 41 | 1 | 4.395 | 0.085245 | 5.7805 | 0.67493 | 9.933 | n.a.
60 | 42 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
61 | A-400x | 1 | 4.395 | 0.85295 | 40.6662 | 7.5149 | 48.0509 | 0.033351
62 | MQ Blank 2 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
63 | 43 | 1 | 4.3992 | 0.47009 | 0.30758 | 2.685 | 0.30983 | n.a.
64 | 44 | 1 | 4.395 | 0.34958 | 0.2269 | 2.0241 | 0.232 | n.a.
65 | 45 | 1 | 4.3909 | 0.35762 | 0.2522 | 2.0919 | 0.25695 | n.a.
66 | 46 | 1 | 4.395 | 0.30243 | 0.33984 | 1.7815 | 0.32808 | n.a.
67 | 47 | 1 | 4.3992 | 0.41799 | 0.39998 | 2.2638 | 0.36232 | n.a.
68 | 48 | 1 | 4.32 | 0.16201 | 4.6086 | 0.56682 | 2.7875 | n.a.
69 | 49 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
70 | 50 | 1 | 4.32 | 0.16831 | 5.4185 | 0.56908 | 3.5002 | n.a.
71 | 51 | 1 | 4.3284 | 0.16142 | 4.9237 | 0.56615 | 3.2064 | n.a.
72 | 52 | 1 | 4.3075 | 0.16562 | 5.0892 | 0.58405 | 3.1518 | n.a.
73 | 53 | 1 | 4.4158 | 0.96723 | 2.246 | 7.9251 | 2.8672 | 0.068953
74 | 54 | 1 | 4.4159 | 1.0187 | 5.7985 | 8.4172 | 7.3162 | 0.084988
75 | 55 | 1 | 4.4075 | 0.67359 | 0.83013 | 4.3005 | 0.86427 | n.a.
76 | 56 | 1 | 4.4075 | 1.0506 | 2.8968 | 8.2652 | 3.5614 | 0.094935
77 | 57 | 1 | 4.4034 | 0.72508 | 2.0997 | 5.2903 | 2.3867 | n.a.
78 | 58 | 1 | 4.3908 | 0.18623 | 20.3056 | 0.71632 | 33.6922 | n.a.
79 | 59 | 1 | 4.3784 | 0.16653 | 26.8423 | 0.61077 | 36.471 | n.a.
80 | 60 | 1 | 4.3617 | 0.28395 | 83.9359 | 0.58829 | 58.0869 | n.a.
81 | 61 | 1 | 4.3908 | 0.17555 | 28.4055 | 0.662 | 37.8715 | n.a.
82 | 62 | 1 | 4.3783 | 0.19879 | 21.8138 | 0.74188 | 20.1315 | n.a.
83 | 63 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
84 | A-200x | 1 | 4.395 | 1.9853 | 41.4484 | 16.2842 | 49.9561 | 0.38612
85 | MQ Blank 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
86 | 64 | 1 | 4.42 | 1.9547 | 3.0264 | 17.1067 | 4.1658 | 0.37659
87 | 65 | 1 | 4.42 | 1.5964 | 2.3142 | 13.776 | 3.167 | 0.26495
88 | 66 | 1 | 4.4117 | 2.3014 | 3.3456 | 20.0434 | 4.5977 | 0.4846
89 | 67 | 1 | 4.4242 | 3.2885 | 4.8636 | 28.7166 | 6.6409 | 0.79211
90 | 68 | 1 | 4.42 | 4.9596 | 7.0079 | 43.5644 | 9.5571 | 1.3127
91 | 69 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
92 | 70 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
93 | 71 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
94 | 72 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
95 | 73 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
96 | 74 | 1 | 4.4367 | 0.24005 | 0.8312 | 1.646 | 0.88063 | n.a.
97 | 75 | 1 | 4.4616 | 0.21328 | 0.68816 | 0.9865 | 0.49278 | n.a.
98 | 76 | 1 | 4.4283 | 0.30914 | 0.73614 | 1.8599 | 0.69157 | n.a.
99 | 77 | 1 | 4.4409 | 0.20119 | 0.71728 | 1.2706 | 0.70177 | n.a.
100 | 78 | 1 | 4.4367 | 0.18256 | 0.46009 | 1.089 | 0.42822 | n.a.
101 | 79 | 1 | 4.3908 | 0.20221 | 9.5908 | 0.88432 | 7.7463 | n.a.
102 | 80 | 1 | 4.4033 | 0.33308 | 18.1526 | 2.0206 | 20.265 | n.a.
103 | 81 | 1 | 4.3991 | 0.24866 | 11.4402 | 1.3741 | 12.9395 | n.a.
104 | 82 | 1 | 4.4034 | 0.233 | 10.9722 | 1.1218 | 12.2688 | n.a.
105 | 83 | 1 | 4.3909 | 0.2043 | 11.1505 | 0.89453 | 11.7418 | n.a.
106 | 84 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
107 | A-800x | 1 | 4.3908 | 0.3338 | 33.7655 | 3.1811 | 43.1943 | n.a.
108 | MQ Blank 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
109 | SB 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
110 | SB 2 | 1 | 4.3117 | 0.02331 | 35.2011 | 0.21624 | 40.3256 | n.a.
111 | SB 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
112 | SB 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
113 | 5 ppm F/10 ppm Anion Std | 1 | 4.4075 | 34.184 | 45.1409 | 273.106 | 51.7363 | 10.4171
| Sum: |  | 412.3724 | 398.6781 | 1171.5306 | 3089.0294 | 1320.8418 | 113.0123
| Average: |  | 4.3869 | 4.2413 | 12.4631 | 32.862 | 14.0515 | 4.7088
| Rel.Std.Dev: |  | 0.8362% | 291.8% | 136.3% | 291.7% | 135.4% | 137.2%
No. | Name | Dilution | Time | Area | Rel.Area | Height | Rel.Height | Amount
|  |  | min | µS*min | % | µS | % | ppm
|  |  | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total
|  |  | Sulfate | Sulfate | Sulfate | Sulfate | Sulfate | Sulfate
1 | DI Water | 1 | 4.8068 | 0.00052081 | 1.5017 | 0.0020495 | 0.93113 | n.a.
2 | 10 ppm Anion Std | 1 | 4.8567 | 2.8683 | 16.3547 | 24.5443 | 15.5785 | 10.3833
3 | DI water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 1600x | 1 | 4.8533 | 0.018706 | 10.2146 | 0.14352 | 10.0943 | n.a.
5 | 800x | 1 | 4.85 | 0.032762 | 12.4411 | 0.26999 | 12.4548 | n.a.
6 | 400x | 1 | 4.8567 | 0.064057 | 14.3436 | 0.55456 | 14.1868 | 0.11243
7 | 200x | 1 | 4.8467 | 0.13908 | 15.4402 | 1.216 | 15.8069 | 0.38721
8 | 20x | 1 | 4.86 | 1.4146 | 16.5347 | 12.1749 | 16.3971 | 5.0589
9 | 5x | 1 | 4.8633 | 5.4904 | 16.9209 | 46.6629 | 17.4713 | 19.9874
10 | 10 ppm Anion Std | 1 | 4.8633 | 2.8676 | 16.3355 | 24.6171 | 15.5977 | 10.3807
11 | 062 | 1 | 4.8933 | 15.3949 | 93.9512 | 126.5257 | 93.1819 | 56.2639 | skip
12 | 055 | 1 | 4.8667 | 1.2104 | 10.11 | 9.4553 | 11.8082 | 4.3112 | skip
13 | 043 | 1 | 4.8567 | 0.28342 | 5.0525 | 2.4319 | 6.3907 | 0.91589 | skip
14 | 10 ppm Anion Std | 1 | 4.86 | 2.776 | 16.3355 | 23.7414 | 15.556 | 10.0454
15 | 10 ppm Anion Std | 1 | 4.8633 | 2.7712 | 16.3074 | 23.7898 | 15.5846 | 10.0277
16 | 1 | 1 | 4.857 | 0.45584 | 4.7395 | 3.9192 | 6.2939 | 1.5474
17 | 2 | 1 | 4.85 | 0.43606 | 4.6678 | 3.7769 | 6.2271 | 1.4749
18 | 3 | 1 | 4.8533 | 0.40558 | 4.472 | 3.6016 | 6.0894 | 1.3633
19 | 4 | 1 | 4.84 | 0.47033 | 4.8127 | 4.0033 | 6.3338 | 1.6005
20 | 5 | 1 | 4.86 | 0.45069 | 4.6582 | 3.9163 | 6.2591 | 1.5285
21 | 6 | 1 | 4.8567 | 0.46235 | 56.3045 | 4.0914 | 56.9176 | 1.5712
22 | 7 | 1 | 4.8533 | 0.44652 | 55.2701 | 3.9468 | 56.4747 | 1.5132
23 | 8 | 1 | 4.85 | 0.42656 | 53.398 | 3.7671 | 54.5636 | 1.4401
24 | 9 | 1 | 4.8467 | 0.46534 | 54.5711 | 4.0988 | 55.4846 | 1.5822
25 | 10 | 1 | 4.85 | 0.46164 | 55.2732 | 4.0501 | 56.092 | 1.5686
26 | 11 | 1 | 4.86 | 0.70102 | 11.9439 | 5.9561 | 14.3211 | 2.4454
27 | 12 | 1 | 4.85 | 0.6348 | 11.8251 | 5.4182 | 14.0426 | 2.2029
28 | 13 | 1 | 4.85 | 0.59214 | 12.6696 | 5.0179 | 14.8794 | 2.0466
29 | 14 | 1 | 4.85 | 0.64861 | 13.9121 | 5.3874 | 16.1086 | 2.2535
30 | 15 | 1 | 4.8533 | 0.66791 | 14.5492 | 5.6268 | 16.8653 | 2.3241
31 | 16 | 1 | 4.8933 | 13.9699 | 93.3733 | 116.1675 | 92.5066 | 51.0445
32 | 17 | 1 | 4.89 | 14.9651 | 93.9744 | 123.6966 | 92.9939 | 54.6897
33 | 18 | 1 | 4.8833 | 13.1778 | 93.7645 | 109.7079 | 92.7074 | 48.1437
34 | 19 | 1 | 4.8867 | 11.9481 | 93.0225 | 100.0079 | 91.9888 | 43.6396
35 | 20 | 1 | 4.8833 | 12.4102 | 93.0756 | 103.5784 | 91.9954 | 45.3322
36 | 21 | 1 | 4.8533 | 0.0046134 | 4.8452 | 0.025082 | 4.1032 | n.a.
37 | A-5x | 1 | 4.8667 | 5.4912 | 16.9462 | 46.5738 | 17.4833 | 19.9901
38 | MQ Blank 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
39 | 22 | 1 | 4.86 | 0.50955 | 4.8836 | 4.558 | 6.8053 | 1.7441
40 | 23 | 1 | 4.85 | 0.5351 | 5.0362 | 4.5076 | 6.675 | 1.8377
41 | 24 | 1 | 4.85 | 0.5245 | 4.9148 | 4.4853 | 6.6263 | 1.7989
42 | 25 | 1 | 4.85 | 0.51993 | 4.9444 | 4.5117 | 6.728 | 1.7821
43 | 26 | 1 | 4.85 | 0.52444 | 5.0202 | 4.4166 | 6.6434 | 1.7987
44 | 27 | 1 | 4.85 | 0.50185 | 60.4709 | 4.43 | 61.3682 | 1.7159
45 | 28 | 1 | 4.8633 | 0.50607 | 60.5649 | 4.4497 | 61.3572 | 1.7314
46 | 29 | 1 | 4.8467 | 0.48606 | 57.9425 | 4.2717 | 59.0567 | 1.6581
47 | 30 | 1 | 4.8533 | 0.53382 | 59.5305 | 4.7115 | 60.4981 | 1.833
48 | 31 | 1 | 4.8533 | 0.5125 | 59.6231 | 4.5071 | 60.426 | 1.7549
49 | 32 | 1 | 4.8633 | 0.80378 | 10.245 | 6.4277 | 11.8574 | 2.8218
50 | 33 | 1 | 4.8567 | 0.78714 | 11.3005 | 6.503 | 13.3626 | 2.7608
51 | 34 | 1 | 4.85 | 0.82234 | 12.0065 | 6.7311 | 14.0746 | 2.8897
52 | 35 | 1 | 4.86 | 1.1213 | 19.4276 | 7.1236 | 18.3891 | 3.9846
53 | 36 | 1 | 4.8567 | 1.1179 | 18.6721 | 7.626 | 18.4457 | 3.9724
54 | 37 | 1 | 4.8933 | 17.4651 | 93.4916 | 142.1829 | 92.772 | 63.8462
55 | 5 ppm F/10 ppm Anion Std | 1 | 4.8567 | 2.7748 | 16.2955 | 23.8669 | 15.5613 | 10.0408
56 | 38 | 1 | 4.8867 | 16.9943 | 93.5487 | 139.1944 | 92.7645 | 62.1219
57 | 39 | 1 | 4.8933 | 17.6629 | 94.2463 | 143.7884 | 93.4734 | 64.5708
58 | 40 | 1 | 4.8967 | 18.2891 | 94.112 | 148.5813 | 93.5044 | 66.8645
59 | 41 | 1 | 4.8967 | 20.0192 | 94.0233 | 161.1195 | 93.5607 | 73.2013
60 | 42 | 1 | 4.8467 | 0.0058567 | 4.4737 | 0.053944 | 5.6684 | n.a.
61 | A-400x | 1 | 4.8533 | 0.063788 | 14.6288 | 0.54931 | 14.4167 | 0.11144
62 | MQ Blank 2 | 1 | 4.8333 | 0.0026922 | 3.1422 | 0.008783 | 1.8468 | n.a.
63 | 43 | 1 | 4.8467 | 0.57456 | 5.3236 | 4.7191 | 6.9154 | 1.9822
64 | 44 | 1 | 4.8467 | 0.5882 | 5.418 | 4.883 | 7.1045 | 2.0322
65 | 45 | 1 | 4.8533 | 0.55151 | 5.4937 | 4.5548 | 7.071 | 1.8978
66 | 46 | 1 | 4.8533 | 0.35713 | 5.6 | 2.9935 | 6.9682 | 1.1859
67 | 47 | 1 | 4.8467 | 0.40866 | 5.5386 | 3.387 | 6.9139 | 1.3746
68 | 48 | 1 | 4.8533 | 0.54515 | 59.2212 | 4.7949 | 59.7211 | 1.8745
69 | 49 | 1 | 4.8567 | 0.51414 | 61.0931 | 4.5684 | 61.4351 | 1.7609
70 | 50 | 1 | 4.8467 | 0.5136 | 61.4254 | 4.5452 | 61.9368 | 1.759
71 | 51 | 1 | 4.8467 | 0.52816 | 61.406 | 4.6747 | 62.0234 | 1.8123
72 | 52 | 1 | 4.8433 | 0.50869 | 59.8603 | 4.4917 | 60.5942 | 1.741
73 | 53 | 1 | 4.8533 | 1.3377 | 25.0303 | 9.22 | 24.1458 | 4.7772
74 | 54 | 1 | 4.86 | 1.2543 | 32.8102 | 9.9244 | 35.2669 | 4.472
75 | 55 | 1 | 4.8567 | 0.79414 | 11.0655 | 6.2181 | 12.4326 | 2.7865
76 | 56 | 1 | 4.86 | 1.2434 | 26.4223 | 9.1279 | 26.7539 | 4.432
77 | 57 | 1 | 4.8567 | 0.73136 | 19.658 | 4.9454 | 18.933 | 2.5565
78 | 58 | 1 | 4.8933 | 21.1724 | 94.7952 | 169.1725 | 94.1217 | 77.4251
79 | 59 | 1 | 4.9033 | 20.9618 | 94.9458 | 167.4095 | 94.2081 | 76.6536
80 | 60 | 1 | 4.8933 | 18.4105 | 94.9627 | 149.3703 | 94.4127 | 67.309
81 | 61 | 1 | 4.8933 | 20.132 | 94.8449 | 161.9007 | 94.1144 | 73.6143
82 | 62 | 1 | 4.8933 | 18.9434 | 94.5425 | 153.3094 | 93.8877 | 69.261
83 | 63 | 1 | 4.8467 | 0.0059318 | 6.216 | 0.035595 | 6.3936 | n.a.
84 | A-200x | 1 | 4.8433 | 0.13887 | 15.659 | 1.2124 | 16.1062 | 0.38643
85 | MQ Blank 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
86 | 64 | 1 | 4.8567 | 0.5307 | 8.9951 | 3.6993 | 9.6557 | 1.8216
87 | 65 | 1 | 4.8567 | 0.58538 | 9.458 | 4.2548 | 10.5182 | 2.0219
88 | 66 | 1 | 4.85 | 0.62404 | 10.3087 | 4.2834 | 10.7901 | 2.1635
89 | 67 | 1 | 4.8533 | 0.51538 | 8.5192 | 3.5428 | 8.9933 | 1.7655
90 | 68 | 1 | 4.8533 | 0.55661 | 8.9803 | 4.0231 | 9.7719 | 1.9165
91 | 69 | 1 | 4.8467 | 0.60194 | 54.1441 | 5.1314 | 53.3892 | 2.0825
92 | 70 | 1 | 4.8467 | 0.58375 | 52.1807 | 4.9515 | 51.4903 | 2.0159
93 | 71 | 1 | 4.85 | 0.5825 | 50.7858 | 4.9512 | 50.3368 | 2.0113
94 | 72 | 1 | 4.85 | 0.58932 | 55.0283 | 5.0278 | 54.9643 | 2.0363
95 | 73 | 1 | 4.85 | 0.59661 | 53.8474 | 5.0918 | 53.0682 | 2.063
96 | 74 | 1 | 4.8633 | 0.82377 | 17.9812 | 5.9141 | 19.682 | 2.895
97 | 75 | 1 | 4.8633 | 0.75392 | 14.3279 | 5.6944 | 16.8844 | 2.6391
98 | 76 | 1 | 4.8567 | 1.023 | 19.1604 | 5.9492 | 16.5391 | 3.6248
99 | 77 | 1 | 4.8567 | 0.74326 | 17.2044 | 5.2646 | 19.0411 | 2.6001
100 | 78 | 1 | 4.8567 | 0.83522 | 15.3552 | 6.0448 | 16.4667 | 2.9369
101 | 79 | 1 | 4.8867 | 15.6764 | 91.5879 | 129.2308 | 91.8697 | 57.2952
102 | 80 | 1 | 4.89 | 17.4289 | 92.3977 | 142.2993 | 92.6099 | 63.7139
103 | 81 | 1 | 4.8933 | 19.78 | 92.536 | 159.6939 | 92.8024 | 72.3251
104 | 82 | 1 | 4.8867 | 16.9103 | 92.5624 | 138.595 | 92.6714 | 61.8143
105 | 83 | 1 | 4.8833 | 15.1512 | 93.0449 | 124.989 | 93.1386 | 55.3714
106 | 84 | 1 | 4.8433 | 0.0029027 | 3.3608 | 0.017302 | 3.1946 | n.a.
107 | A-800x | 1 | 4.8467 | 0.032021 | 12.9604 | 0.26764 | 12.8748 | n.a.
108 | MQ Blank 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
109 | SB 1 | 1 | 4.8367 | 0.0046565 | 3.8398 | 0.029027 | 3.5866 | n.a.
110 | SB 2 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
111 | SB 3 | 1 | 4.8367 | 0.0026395 | 3.5048 | 0.0090587 | 2.2822 | n.a.
112 | SB 4 | 1 | 4.8333 | 0.0023869 | 3.223 | 0.0076386 | 1.9603 | n.a.
113 | 5 ppm F/10 ppm Anion Std | 1 | 4.85 | 2.7744 | 16.3013 | 23.8895 | 15.5697 | 10.0394
| Sum: |  | 524.8472 | 426.0352 | 3859.3851 | 3479.3809 | 3903.2649 | 1548.2645
| Average: |  | 4.8597 | 3.9448 | 35.735 | 32.2165 | 36.1413 | 16.1278
| Rel.Std.Dev: |  | 0.3585% | 167.5% | 94.83% | 166.9% | 92.53% | 155.7%
No. | Name | Dilution | Time | Area | Rel.Area | Height | Rel.Height | Amount
|  |  | min | µS*min | % | µS | % | ppm
|  |  | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total
|  |  | Bromide | Bromide | Bromide | Bromide | Bromide | Bromide
1 | DI Water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 10 ppm Anion Std | 1 | 6.2767 | 1.6685 | 9.5139 | 13.0248 | 8.267 | 10.3999
3 | DI water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 1600x | 1 | 6.2933 | 0.0078155 | 4.2678 | 0.062315 | 4.3828 | n.a.
5 | 800x | 1 | 6.29 | 0.016155 | 6.1348 | 0.1295 | 5.9737 | n.a.
6 | 400x | 1 | 6.2967 | 0.034455 | 7.7151 | 0.27579 | 7.0552 | 0.11375
7 | 200x | 1 | 6.29 | 0.077179 | 8.5682 | 0.61674 | 8.0169 | 0.38268
8 | 20x | 1 | 6.2867 | 0.82063 | 9.5923 | 6.4099 | 8.6328 | 5.0625
9 | 5x | 1 | 6.26 | 3.1915 | 9.8358 | 24.942 | 9.3387 | 19.9864
10 | 10 ppm Anion Std | 1 | 6.2833 | 1.6674 | 9.4985 | 13.0214 | 8.2505 | 10.3927
11 | 062 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | skip
12 | 055 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | skip
13 | 043 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | skip
14 | 10 ppm Anion Std | 1 | 6.2733 | 1.6184 | 9.5235 | 12.6504 | 8.2889 | 10.0843
15 | 10 ppm Anion Std | 1 | 6.2767 | 1.6155 | 9.5064 | 12.6356 | 8.2775 | 10.0657
16 | 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
17 | 2 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
18 | 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
19 | 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
20 | 5 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
21 | 6 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
22 | 7 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
23 | 8 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
24 | 9 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
25 | 10 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
26 | 11 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
27 | 12 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
28 | 13 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
29 | 14 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
30 | 15 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
31 | 16 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
32 | 17 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
33 | 18 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
34 | 19 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
35 | 20 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
36 | 21 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
37 | A-5x | 1 | 6.26 | 3.1962 | 9.8636 | 24.9759 | 9.3757 | 20.0159
38 | MQ Blank 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
39 | 22 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
40 | 23 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
41 | 24 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
42 | 25 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
43 | 26 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
44 | 27 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
45 | 28 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
46 | 29 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
47 | 30 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
48 | 31 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
49 | 32 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
50 | 33 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
51 | 34 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
52 | 35 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
53 | 36 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
54 | 37 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
55 | 5 ppm F/10 ppm Anion Std | 1 | 6.27 | 1.6193 | 9.5095 | 12.6732 | 8.263 | 10.0897
56 | 38 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
57 | 39 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
58 | 40 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
59 | 41 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
60 | 42 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
61 | A-400x | 1 | 6.29 | 0.033964 | 7.7892 | 0.27391 | 7.1888 | 0.11066
62 | MQ Blank 2 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
63 | 43 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
64 | 44 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
65 | 45 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
66 | 46 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
67 | 47 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
68 | 48 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
69 | 49 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
70 | 50 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
71 | 51 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
72 | 52 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
73 | 53 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
74 | 54 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
75 | 55 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
76 | 56 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
77 | 57 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
78 | 58 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
79 | 59 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
80 | 60 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
81 | 61 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
82 | 62 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
83 | 63 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
84 | A-200x | 1 | 6.28 | 0.076755 | 8.6552 | 0.61667 | 8.1925 | 0.38001
85 | MQ Blank 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
86 | 64 | 1 | 6.27 | 0.0067667 | 0.11469 | 0.046165 | 0.1205 | n.a.
87 | 65 | 1 | 6.2667 | 0.0061779 | 0.099817 | 0.042198 | 0.10432 | n.a.
88 | 66 | 1 | 6.27 | 0.0056679 | 0.09363 | 0.040384 | 0.10173 | n.a.
89 | 67 | 1 | 6.2633 | 0.006639 | 0.10974 | 0.046074 | 0.11696 | n.a.
90 | 68 | 1 | 6.2633 | 0.0062098 | 0.10019 | 0.042711 | 0.10374 | n.a.
91 | 69 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
92 | 70 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
93 | 71 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
94 | 72 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
95 | 73 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
96 | 74 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
97 | 75 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
98 | 76 | 1 | 6.2767 | 0.0091282 | 0.17096 | 0.058063 | 0.16142 | n.a.
99 | 77 | 1 | 6.2733 | 0.010389 | 0.24048 | 0.062149 | 0.22478 | n.a.
100 | 78 | 1 | 6.2767 | 0.0086983 | 0.15991 | 0.070905 | 0.19316 | n.a.
101 | 79 | 1 | 6.2667 | 0.0063711 | 0.037222 | 0.033011 | 0.023467 | n.a.
102 | 80 | 1 | 6.2633 | 0.0075866 | 0.040219 | 0.038204 | 0.024863 | n.a.
103 | 81 | 1 | 6.2633 | 0.0088375 | 0.041344 | 0.044889 | 0.026086 | n.a.
104 | 82 | 1 | 6.2567 | 0.0060955 | 0.033365 | 0.031336 | 0.020953 | n.a.
105 | 83 | 1 | 6.26 | 0.0055529 | 0.034101 | 0.028487 | 0.021227 | n.a.
106 | 84 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
107 | A-800x | 1 | 6.2833 | 0.016005 | 6.478 | 0.12888 | 6.1996 | n.a.
108 | MQ Blank 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
109 | SB 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
110 | SB 2 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
111 | SB 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
112 | SB 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
113 | 5 ppm F/10 ppm Anion Std | 1 | 6.2667 | 1.6153 | 9.4909 | 12.6623 | 8.2525 | 10.0647
| Sum: |  | 181.9467 | 17.3691 | 137.2184 | 135.6837 | 125.1992 | 107.149
| Average: |  | 6.274 | 0.59894 | 4.7317 | 4.6787 | 4.3172 | 8.2422
| Rel.Std.Dev: |  | 0.1809% | 163.1% | 93.29% | 163.2% | 92.39% | 83.26%
No. | Name | Dilution | Time | Area | Rel.Area | Height | Rel.Height | Amount
|  |  | min | mAU*min | % | mAU | % | ppm
|  |  | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1 | UV_VIS_1
|  |  | Nitrate UV | Nitrate UV | Nitrate UV | Nitrate UV | Nitrate UV | Nitrate UV
1 | DI Water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 10 ppm Anion Std | 1 | 6.686 | 35.1915 | 45.5553 | 220.8968 | 41.5414 | 10.3403
3 | DI water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 1600x | 1 | 6.7068 | 0.45449 | 58.3405 | 3.0428 | 53.4701 | n.a.
5 | 800x | 1 | 6.7151 | 0.60324 | 49.3045 | 3.9206 | 44.1452 | n.a.
6 | 400x | 1 | 6.711 | 1.0103 | 42.1587 | 6.4824 | 38.2207 | 0.08372
7 | 200x | 1 | 6.711 | 2.0331 | 41.008 | 12.5803 | 37.1934 | 0.39061
8 | 20x | 1 | 6.6902 | 17.9916 | 43.8653 | 113.9728 | 40.3112 | 5.1792
9 | 5x | 1 | 6.6569 | 67.2436 | 46.6865 | 409.1535 | 42.7479 | 19.9579
10 | 10 ppm Anion Std | 1 | 6.6777 | 35.2409 | 45.6641 | 220.8397 | 41.4856 | 10.3551
11 | 062 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | skip
12 | 055 | 1 | 6.6194 | 137.8662 | 98.8057 | 792.9225 | 98.8035 | 41.1493 | skip
13 | 043 | 1 | 6.6485 | 79.0155 | 99.2346 | 484.9723 | 99.5372 | 23.4903 | skip
14 | 10 ppm Anion Std | 1 | 6.6818 | 35.2913 | 46.2418 | 222.0546 | 42.132 | 10.3702
15 | 10 ppm Anion Std | 1 | 6.686 | 35.2578 | 46.297 | 221.7229 | 42.0845 | 10.3602
16 | 1 | 1 | 6.6159 | 140.7465 | 99.3841 | 807.5825 | 99.7042 | 42.0135
17 | 2 | 1 | 6.6242 | 136.2893 | 99.2631 | 785.5941 | 99.6266 | 40.6761
18 | 3 | 1 | 6.62 | 133.2092 | 99.176 | 770.2305 | 99.5159 | 39.7519
19 | 4 | 1 | 6.6158 | 142.5632 | 99.2818 | 816.8778 | 99.6245 | 42.5587
20 | 5 | 1 | 6.6158 | 141.6042 | 99.2938 | 811.742 | 99.5557 | 42.2709
21 | 6 | 1 | 6.7033 | 2.5232 | 74.999 | 16.7487 | 86.1276 | 0.53767
22 | 7 | 1 | 6.7034 | 2.7946 | 76.6533 | 18.4602 | 87.5139 | 0.61913
23 | 8 | 1 | 6.7033 | 3.16 | 77.1735 | 19.2657 | 88.3256 | 0.72877
24 | 9 | 1 | 6.7033 | 2.9782 | 76.3594 | 19.7691 | 87.0663 | 0.67422
25 | 10 | 1 | 6.7075 | 2.8411 | 76.4797 | 17.5807 | 87.2835 | 0.63307
26 | 11 | 1 | 6.6575 | 70.9024 | 98.7348 | 438.5859 | 99.4226 | 21.0558
27 | 12 | 1 | 6.6575 | 63.9603 | 98.9798 | 400.8239 | 99.4493 | 18.9728
28 | 13 | 1 | 6.67 | 53.4478 | 98.098 | 337.2131 | 99.0694 | 15.8183
29 | 14 | 1 | 6.67 | 49.3204 | 98.0385 | 313.9943 | 98.9699 | 14.5799
30 | 15 | 1 | 6.67 | 49.1793 | 98.3043 | 311.9294 | 99.2292 | 14.5375
31 | 16 | 1 | 6.7117 | 0.50103 | 34.7587 | 2.8119 | 57.7339 | n.a.
32 | 17 | 1 | 6.7033 | 0.93936 | 52.6333 | 3.5252 | 62.632 | 0.062431
33 | 18 | 1 | 6.7033 | 0.80659 | 49.239 | 3.7265 | 64.8775 | 0.022592
34 | 19 | 1 | 6.7075 | 1.8457 | 63.7527 | 10.903 | 84.7433 | 0.33438
35 | 20 | 1 | 6.7116 | 1.4171 | 58.3805 | 7.5824 | 78.9169 | 0.20577
36 | 21 | 1 | 6.7033 | 0.13224 | 100 | 0.88152 | 100 | n.a.
37 | A-5x | 1 | 6.6533 | 71.9845 | 49.1513 | 422.6762 | 43.6682 | 21.3805
38 | MQ Blank 1 | 1 | 6.7075 | 0.066924 | 100 | 0.44508 | 100 | n.a.
39 | 22 | 1 | 6.6116 | 149.3948 | 99.4439 | 851.8698 | 99.4727 | 44.6086
40 | 23 | 1 | 6.6159 | 151.919 | 99.3305 | 862.3628 | 99.6176 | 45.366
41 | 24 | 1 | 6.6158 | 152.1934 | 99.1867 | 864.1359 | 99.4198 | 45.4484
42 | 25 | 1 | 6.6158 | 150.6249 | 99.2575 | 855.1406 | 99.4844 | 44.9777
43 | 26 | 1 | 6.62 | 149.1309 | 99.2217 | 848.864 | 99.4492 | 44.5294
44 | 27 | 1 | 6.7033 | 2.3649 | 72.4269 | 15.7207 | 83.9067 | 0.49019
45 | 28 | 1 | 6.6992 | 2.2726 | 70.6926 | 15.0607 | 82.0171 | 0.4625
46 | 29 | 1 | 6.7075 | 2.5849 | 74.5696 | 17.1546 | 84.5934 | 0.55621
47 | 30 | 1 | 6.7033 | 3.0567 | 73.8228 | 18.8127 | 84.7578 | 0.69778
48 | 31 | 1 | 6.7034 | 2.9309 | 72.5127 | 17.2916 | 85.068 | 0.66003
49 | 32 | 1 | 6.6408 | 92.7444 | 98.5567 | 560.2929 | 99.2059 | 27.6098
50 | 33 | 1 | 6.6492 | 78.942 | 98.7438 | 486.1119 | 99.083 | 23.4682
51 | 34 | 1 | 6.6533 | 74.8514 | 97.5477 | 463.1527 | 97.4637 | 22.2408
52 | 35 | 1 | 6.6659 | 51.7329 | 97.5483 | 328.7046 | 97.5173 | 15.3038
53 | 36 | 1 | 6.6617 | 57.03 | 97.1761 | 360.0174 | 97.1179 | 16.8932
54 | 37 | 1 | 6.7033 | 3.0177 | 74.399 | 18.1608 | 90.6077 | 0.68606
55 | 5 ppm F/10 ppm Anion Std | 1 | 6.6742 | 35.2659 | 46.4721 | 222.2431 | 42.0446 | 10.3626
56 | 38 | 1 | 6.6992 | 2.6139 | 73.2089 | 15.8951 | 90.2517 | 0.56489
57 | 39 | 1 | 6.7034 | 1.0332 | 57.0834 | 5.2549 | 71.6682 | 0.090603
58 | 40 | 1 | 6.7033 | 1.0594 | 56.3185 | 5.0575 | 66.0921 | 0.09846
59 | 41 | 1 | 6.7075 | 0.92887 | 62.9877 | 5.2053 | 76.6071 | 0.059283
60 | 42 | 1 | 6.7033 | 0.29273 | 100 | 1.973 | 100 | n.a.
61 | A-400x | 1 | 6.6991 | 0.94593 | 45.099 | 6.0769 | 38.8562 | 0.064401
62 | MQ Blank 2 | 1 | 6.7075 | 0.066609 | 100 | 0.46315 | 100 | n.a.
63 | 43 | 1 | 6.6117 | 151.7512 | 99.2908 | 862.0027 | 99.4694 | 45.3157
64 | 44 | 1 | 6.6117 | 153.0596 | 99.3436 | 868.3309 | 99.5266 | 45.7083
65 | 45 | 1 | 6.6117 | 140.6184 | 99.1668 | 809.64 | 99.4503 | 41.9751
66 | 46 | 1 | 6.6409 | 88.3685 | 99.2993 | 539.5808 | 99.3703 | 26.2968
67 | 47 | 1 | 6.6366 | 103.5363 | 99.0762 | 620.9448 | 99.383 | 30.8481
68 | 48 | 1 | 6.6991 | 2.6237 | 74.634 | 17.4387 | 85.7587 | 0.56785
69 | 49 | 1 | 6.6992 | 2.4332 | 82.3938 | 16.0766 | 92.2951 | 0.51067
70 | 50 | 1 | 6.7033 | 2.2179 | 71.4016 | 13.8706 | 85.314 | 0.44607
71 | 51 | 1 | 6.7034 | 2.4396 | 74.4142 | 15.2958 | 86.6295 | 0.5126
72 | 52 | 1 | 6.6991 | 2.3906 | 73.4593 | 15.8226 | 85.3864 | 0.4979
73 | 53 | 1 | 6.6741 | 41.4657 | 96.2887 | 266.3219 | 96.3515 | 12.2229
74 | 54 | 1 | 6.6908 | 15.8741 | 90.3563 | 104.2325 | 90.598 | 4.5438
75 | 55 | 1 | 6.645 | 79.9186 | 98.4912 | 491.6411 | 98.8061 | 23.7613
76 | 56 | 1 | 6.6742 | 34.3747 | 94.777 | 221.5597 | 95.4673 | 10.0952
77 | 57 | 1 | 6.6784 | 33.1404 | 95.9675 | 214.309 | 96.6845 | 9.7248
78 | 58 | 1 | 6.6991 | 0.089995 | 9.8126 | 0.31441 | 14.7883 | n.a.
79 | 59 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
80 | 60 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
81 | 61 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
82 | 62 | 1 | 6.695 | 0.26591 | 29.1796 | 1.7886 | 48.5353 | n.a.
83 | 63 | 1 | 6.6992 | 0.32312 | 100 | 2.1846 | 100 | n.a.
84 | A-200x | 1 | 6.6991 | 1.9677 | 41.0799 | 12.4005 | 38.0417 | 0.37099
85 | MQ Blank 3 | 1 | 6.6992 | 0.07442 | 100 | 0.4841 | 100 | n.a.
86 | 64 | 1 | 6.6533 | 62.47 | 96.7176 | 392.6398 | 95.6153 | 18.5256
87 | 65 | 1 | 6.6533 | 67.2151 | 97.4417 | 420.3059 | 96.6245 | 19.9494
88 | 66 | 1 | 6.6533 | 66.2617 | 96.3249 | 414.9322 | 95.1794 | 19.6633
89 | 67 | 1 | 6.6575 | 64.2244 | 94.9853 | 403.2163 | 93.2472 | 19.052
90 | 68 | 1 | 6.6533 | 65.6956 | 92.827 | 411.743 | 90.3274 | 19.4935
91 | 69 | 1 | 6.7033 | 1.3289 | 82.7763 | 8.8167 | 86.7048 | 0.17932
92 | 70 | 1 | 6.6992 | 1.4917 | 84.2501 | 9.9109 | 87.8764 | 0.22818
93 | 71 | 1 | 6.695 | 1.838 | 86.774 | 12.2038 | 89.9922 | 0.33207
94 | 72 | 1 | 6.6949 | 1.702 | 86.0074 | 11.3489 | 89.49 | 0.29127
95 | 73 | 1 | 6.6992 | 1.372 | 83.0108 | 9.1555 | 86.9862 | 0.19226
96 | 74 | 1 | 6.6783 | 28.1315 | 97.4085 | 183.2207 | 98.0279 | 8.2218
97 | 75 | 1 | 6.6783 | 30.5099 | 98.4447 | 197.6072 | 98.7095 | 8.9355
98 | 76 | 1 | 6.67 | 41.1699 | 98.0367 | 264.9666 | 98.5235 | 12.1342
99 | 77 | 1 | 6.6867 | 27.3417 | 97.4759 | 178.0323 | 98.3272 | 7.9848
100 | 78 | 1 | 6.67 | 38.9712 | 98.215 | 251.1742 | 98.767 | 11.4744
101 | 79 | 1 | 6.7033 | 1.4207 | 67.3825 | 9.4622 | 82.8856 | 0.20686
102 | 80 | 1 | 6.7033 | 1.0324 | 56.2647 | 6.8863 | 69.0633 | 0.090344
103 | 81 | 1 | 6.7033 | 1.2851 | 59.123 | 8.1773 | 77.0012 | 0.16616
104 | 82 | 1 | 6.7033 | 1.2279 | 57.8229 | 6.9322 | 75.8151 | 0.14901
105 | 83 | 1 | 6.695 | 1.1092 | 60.5367 | 5.6565 | 74.2491 | 0.11339
106 | 84 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
107 | A-800x | 1 | 6.6908 | 0.5219 | 52.7932 | 3.3855 | 45.9694 | n.a.
108 | MQ Blank 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
109 | SB 1 | 1 | 6.695 | 0.39708 | 100 | 2.6751 | 100 | n.a.
110 | SB 2 | 1 | 6.7075 | 0.04291 | 64.7989 | 0.32 | 59.6744 | n.a.
111 | SB 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
112 | SB 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
113 | 5 ppm F/10 ppm Anion Std | 1 | 6.67 | 35.2329 | 46.5259 | 221.9537 | 42.0462 | 10.3527
| Sum: |  | 687.7845 | 4086.3378 | 8091.2787 | 24425.4963 | 8471.9898 | 1205.4837
| Average: |  | 6.6775 | 39.6732 | 78.5561 | 237.1407 | 82.2523 | 13.5448
| Rel.Std.Dev: |  | 0.4713% | 126.3% | 28.36% | 122.3% | 26.3% | 113.5%
No. | Name | Dilution | Time | Area | Rel.Area | Height | Rel.Height | Amount
|  |  | min | µS*min | % | µS | % | ppm
|  |  | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total | CD_1_Total
|  |  | Phosphate | Phosphate | Phosphate | Phosphate | Phosphate | Phosphate
1 | DI Water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 10 ppm Anion Std | 1 | 8.99 | 1.216 | 6.9333 | 5.6564 | 3.5902 | 10.1539
3 | DI water | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 1600x | 1 | 9.05 | 0.011586 | 6.3268 | 0.050966 | 3.5846 | 0.10948
5 | 800x | 1 | 9.0467 | 0.024311 | 9.2318 | 0.10856 | 5.0079 | 0.2156
6 | 400x | 1 | 9.0533 | 0.050054 | 11.208 | 0.22942 | 5.869 | 0.43029
7 | 200x | 1 | 9.0367 | 0.11486 | 12.751 | 0.53306 | 6.9292 | 0.97072
8 | 20x | 1 | 8.9933 | 1.2218 | 14.2817 | 5.6326 | 7.586 | 10.2027
9 | 5x | 1 | 8.8667 | 4.7888 | 14.7587 | 19.1049 | 7.1532 | 39.9512
10 | 10 ppm Anion Std | 1 | 9 | 1.2293 | 7.003 | 5.6936 | 3.6075 | 10.2652
11 | 062 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | skip
12 | 055 | 1 | 9.0733 | 0.055274 | 0.46168 | 0.23375 | 0.29192 | 0.47383 | skip
13 | 043 | 1 | 9.07 | 0.011681 | 0.20823 | 0.046688 | 0.12269 | 0.11026 | skip
14 | 10 ppm Anion Std | 1 | 8.9933 | 1.1638 | 6.8485 | 5.415 | 3.5481 | 9.7189
15 | 10 ppm Anion Std | 1 | 8.9967 | 1.1799 | 6.9435 | 5.4513 | 3.5712 | 9.8534
16 | 1 | 1 | 9.0703 | 0.016311 | 0.16959 | 0.056617 | 0.090921 | 0.14888
17 | 2 | 1 | 9.0633 | 0.014785 | 0.15827 | 0.043897 | 0.072375 | 0.13615
18 | 3 | 1 | 9.07 | 0.011948 | 0.13174 | 0.041749 | 0.070588 | 0.11249
19 | 4 | 1 | 9.05 | 0.01563 | 0.15993 | 0.058493 | 0.092544 | 0.1432
20 | 5 | 1 | 9.07 | 0.016422 | 0.16973 | 0.060561 | 0.09679 | 0.14981
21 | 6 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
22 | 7 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
23 | 8 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
24 | 9 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
25 | 10 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
26 | 11 | 1 | 9.0633 | 0.027827 | 0.47411 | 0.095951 | 0.23071 | 0.24492
27 | 12 | 1 | 9.06 | 0.022458 | 0.41835 | 0.082413 | 0.21359 | 0.20015
28 | 13 | 1 | 9.0533 | 0.037824 | 0.80929 | 0.1432 | 0.42462 | 0.32829
29 | 14 | 1 | 9.0533 | 0.029595 | 0.63479 | 0.10506 | 0.31412 | 0.25967
30 | 15 | 1 | 9.0533 | 0.034091 | 0.74261 | 0.12879 | 0.38602 | 0.29716
31 | 16 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
32 | 17 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
33 | 18 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
34 | 19 | 1 | 9.0967 | 0.0027996 | 0.021797 | 0.0084339 | 0.0077577 | 0.036196
35 | 20 | 1 | 9.0967 | 0.0022975 | 0.017231 | 0.0068608 | 0.0060935 | 0.032008
36 | 21 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
37 | A-5x | 1 | 8.8667 | 4.749 | 14.6558 | 18.986 | 7.1272 | 39.6191
38 | MQ Blank 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
39 | 22 | 1 | 9.0633 | 0.022772 | 0.21825 | 0.093099 | 0.139 | 0.20276
40 | 23 | 1 | 9.0567 | 0.021945 | 0.20654 | 0.083384 | 0.12348 | 0.19587
41 | 24 | 1 | 9.06 | 0.023758 | 0.22263 | 0.087088 | 0.12866 | 0.21099
42 | 25 | 1 | 9.0567 | 0.018888 | 0.17961 | 0.073885 | 0.11018 | 0.17037
43 | 26 | 1 | 9.0567 | 0.024883 | 0.23819 | 0.086042 | 0.12942 | 0.22037
44 | 27 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
45 | 28 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
46 | 29 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
47 | 30 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
48 | 31 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
49 | 32 | 1 | 9.0633 | 0.037957 | 0.4838 | 0.16306 | 0.3008 | 0.3294
50 | 33 | 1 | 9.06 | 0.049464 | 0.71013 | 0.21634 | 0.44455 | 0.42537
51 | 34 | 1 | 9.0467 | 0.05238 | 0.76478 | 0.2282 | 0.47716 | 0.44969
52 | 35 | 1 | 9.0633 | 0.031487 | 0.54557 | 0.12466 | 0.3218 | 0.27545
53 | 36 | 1 | 9.06 | 0.037455 | 0.62558 | 0.15498 | 0.37486 | 0.32521
54 | 37 | 1 | 9.11 | 0.00312 | 0.016702 | 0.0093806 | 0.0061207 | 0.038868
55 | 5 ppm F/10 ppm Anion Std | 1 | 8.9867 | 1.176 | 6.9065 | 5.4679 | 3.5651 | 9.8207
56 | 38 | 1 | 9.0967 | 0.002653 | 0.014604 | 0.0094944 | 0.0063274 | 0.034973
57 | 39 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
58 | 40 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
59 | 41 | 1 | 9.1233 | 0.0025845 | 0.012139 | 0.0073762 | 0.0042833 | 0.034402
60 | 42 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
61 | A-400x | 1 | 9.05 | 0.046858 | 10.7462 | 0.20492 | 5.3782 | 0.40364
62 | MQ Blank 2 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
63 | 43 | 1 | 9.0567 | 0.023632 | 0.21897 | 0.077037 | 0.11289 | 0.20994
64 | 44 | 1 | 9.0533 | 0.026111 | 0.24051 | 0.089881 | 0.13077 | 0.23061
65 | 45 | 1 | 9.0567 | 0.022981 | 0.22892 | 0.087172 | 0.13533 | 0.2045
66 | 46 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
67 | 47 | 1 | 9.0433 | 0.015755 | 0.21353 | 0.06004 | 0.12256 | 0.14425
68 | 48 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
69 | 49 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
70 | 50 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
71 | 51 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
72 | 52 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
73 | 53 | 1 | 9.05 | 0.027417 | 0.51302 | 0.11385 | 0.29815 | 0.2415
74 | 54 | 1 | 9.05 | 0.020625 | 0.53948 | 0.084549 | 0.30045 | 0.18485
75 | 55 | 1 | 9.0533 | 0.046739 | 0.65126 | 0.2005 | 0.40087 | 0.40265
76 | 56 | 1 | 9.0567 | 0.033777 | 0.71775 | 0.14363 | 0.42097 | 0.29454
77 | 57 | 1 | 9.0533 | 0.03909 | 1.0507 | 0.17033 | 0.6521 | 0.33885
78 | 58 | 1 | 9.1067 | 0.0023627 | 0.010578 | 0.0065419 | 0.0036397 | 0.032552
79 | 59 | 1 | 9.13 | 0.0023839 | 0.010798 | 0.0061765 | 0.0034757 | 0.032729
80 | 60 | 1 | 9.1067 | 0.0030082 | 0.015517 | 0.008909 | 0.0056311 | 0.037936
81 | 61 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
82 | 62 | 1 | 9.1033 | 0.0032473 | 0.016207 | 0.0096988 | 0.0059396 | 0.03993
83 | 63 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
84 | A-200x | 1 | 9.0233 | 0.11091 | 12.5065 | 0.50368 | 6.6914 | 0.93781
85 | MQ Blank 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
86 | 64 | 1 | 9.0567 | 0.022866 | 0.38757 | 0.068991 | 0.18008 | 0.20355
87 | 65 | 1 | 9.06 | 0.020911 | 0.33786 | 0.061443 | 0.15189 | 0.18724
88 | 66 | 1 | 9.05 | 0.016439 | 0.27156 | 0.044169 | 0.11126 | 0.14995
89 | 67 | 1 | 9.0533 | 0.017142 | 0.28335 | 0.050666 | 0.12861 | 0.15581
90 | 68 | 1 | 9.0533 | 0.023067 | 0.37216 | 0.070083 | 0.17023 | 0.20523
91 | 69 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
92 | 70 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
93 | 71 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
94 | 72 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
95 | 73 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
96 | 74 | 1 | 9.06 | 0.048934 | 1.0681 | 0.20019 | 0.66623 | 0.42095
97 | 75 | 1 | 9.05 | 0.081735 | 1.5533 | 0.32173 | 0.95397 | 0.6945
98 | 76 | 1 | 9.0467 | 0.061841 | 1.1582 | 0.2519 | 0.70029 | 0.52859
99 | 77 | 1 | 9.0467 | 0.072415 | 1.6762 | 0.30044 | 1.0867 | 0.61678
100 | 78 | 1 | 9.0533 | 0.053347 | 0.98076 | 0.21273 | 0.57949 | 0.45775
101 | 79 | 1 | 9.0767 | 0.0061394 | 0.035869 | 0.016353 | 0.011625 | 0.06405
102 | 80 | 1 | 9.0867 | 0.003627 | 0.019228 | 0.010415 | 0.0067784 | 0.043097
103 | 81 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
104 | 82 | 1 | 9.0767 | 0.0069224 | 0.037891 | 0.022979 | 0.015365 | 0.070579
105 | 83 | 1 | 9.07 | 0.0046041 | 0.028274 | 0.014535 | 0.010831 | 0.051245
106 | 84 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
107 | A-800x | 1 | 9.03 | 0.023257 | 9.4132 | 0.10076 | 4.8467 | 0.20681
108 | MQ Blank 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
109 | SB 1 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
110 | SB 2 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
111 | SB 3 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
112 | SB 4 | 1 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
113 | 5 ppm F/10 ppm Anion Std | 1 | 8.97 | 1.1832 | 6.9523 | 5.4669 | 3.563 | 9.8809
| Sum: |  | 633.5737 | 19.6252 | 178.9204 | 83.7605 | 93.9719 | 164.5713
| Average: |  | 9.0511 | 0.28036 | 2.556 | 1.1966 | 1.3425 | 2.351
| Rel.Std.Dev: |  | 0.4845% | 303.9% | 167.1% | 291.9% | 163.8% | 302.2%

## Audit Trail
16 | 1 |  |  | Audit Trail
Sample Name: |  | 1 | Injection Volume: |  | 10
Vial Position: |  | GA1 | Channel: |  | UV_VIS_1
Injection Type: |  | Unknown | Wavelength: |  | 215
Instrument Method: |  | AS11HC_SM_UV | Bandwidth: |  | n.a.
Processing Method: |  | ANION_Quant with UV | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-19 12:29 | Sample Weight: |  | 1
Run Time (min): |  | 14.9957 | Int STD Amount: |  | 1
Day Time | Ret. Time | Command/Message
(UTC-07:00) | min
12:22:53 |  | Audit trail of injection "chrom://desktop-5bv84qa/ChromeleonLocal/Instrument Data/SYSTEM_1/60511 Compass/091522.seq/150.smp". Executed 9/19/2022 on instrument SYSTEM_1 (instrument controller DESKTOP-5BV84QA).
12:22:53 |  | Start of injection 20 "1", using instrument method "AS11HC_SM_UV".
12:22:53 | 0.000 | Entered stage "Instrument Setup"
12:22:53 | 0.000 | DC.Compartment_TC.AcquireExclusiveAccess
12:22:53 | 0.000 | Instrument SYSTEM_1 waits for exclusive access.
12:26:52 | 0.000 | DC.Column_TC.AcquireExclusiveAccess
12:26:52 | 0.000 | Exclusive access granted to instrument SYSTEM_1.
12:26:52 | 0.000 | Exclusive access granted to instrument SYSTEM_1.
12:26:52 | 0.000 | Wait Sampler.Ready
12:26:52 | 0.000 | Wait finished
12:26:52 | 0.000 | DC.Suppressor1.VoltageSet = 4.8 V
12:26:52 | 0.000 | Sampler.DiverterValve = Position_1
12:26:52 | 0.000 | Wait Sampler.Ready
12:26:52 | 0.000 | Wait finished
12:26:52 | 0.000 | Pump_1.Flow = 0.300 ml/min
12:26:52 | 0.000 | Sampler.Temperature.Nominal = 15.0 °C
12:26:52 | 0.000 | Log TempCtrl: On
12:26:52 | 0.000 | Sampler.CycleTime = 0 min
12:26:52 | 0.000 | EluentGenerator.EGC_1.Concentration = 30.00 mM
12:26:52 | 0.000 | UV.TimeConstant = 0.12 s
12:26:52 | 0.000 | Pump_1.Pressure.UpperLimit = 5000 psi
12:26:52 | 0.000 | Pump_1.%A.Equate = "%A"
12:26:52 | 0.000 | Pump_1.Pressure.LowerLimit = 200 psi
12:26:52 | 0.000 | Pump_1.%B.Equate = "%B"
12:26:52 | 0.000 | Pump_1.%C.Equate = "%C"
12:26:52 | 0.000 | Pump_1.%D.Equate = "%D"
12:26:52 | 0.000 | DC.Column_TC.Mode = On
12:26:52 | 0.000 | DC.Column_TC.TemperatureSet = 30.00 °C
12:26:52 | 0.000 | DC.Compartment_TC.Mode = On
12:26:52 | 0.000 | DC.Compartment_TC.TemperatureSet = 20.00 °C
12:26:52 | 0.000 | DC.Suppressor1.Type = ADRS_2mm
12:26:52 | 0.000 | CDet1.CellHeater.Mode = On
12:26:52 | 0.000 | CDet1.CellHeater.TemperatureSet = 35.00 °C
12:26:52 | 0.000 | UV.UV_VIS_1.Wavelength = 215 nm
12:26:52 | 0.000 | Pump_1.Curve = 5
12:26:52 | 0.000 | Pump_1.%B.Value = 0.0 %
12:26:52 | 0.000 | Pump_1.%C.Value = 0.0 %
12:26:52 | 0.000 | Pump_1.%D.Value = 0.0 %
12:26:52 | 0.000 | UV.UV_VIS_2.Wavelength = 210 nm
12:26:52 | 0.000 | Entered stage "Inject Preparation"
12:26:52 | 0.000 | Wait UV.Ready
12:26:52 | 0.000 | Wait finished
12:26:52 | 0.000 | UV.Autozero
12:26:52 | 0.000 | Wait UV.Ready
12:26:57 | 0.000 | Wait finished
12:26:57 | 0.000 | Entered stage "Inject"
12:26:57 | 0.000 | Wait Sampler.CycleTimeState
12:26:57 | 0.000 | Wait finished
12:26:57 | 0.000 | Sampler.Inject
12:26:57 | 0.000 | Injecting from vial position GA1.
12:26:57 | 0.000 | Injection Volume is 10.0 µl.
12:26:57 | 0.000 | Waiting for inject response on AsapInjectState1.
12:29:02 | 0.000 | Got inject response.
12:29:02 | 0.000 | Entered stage "Start Run"
12:29:02 | 0.000 | UV.UV_VIS_1.AcqOn
12:29:02 | 0.000 | CDet1.CD_1.AcqOn
12:29:02 | 0.000 | CDet1.CD_1_Total.AcqOn
12:29:02 | 0.000 | CDet1.Autozero
12:29:02 | 0.000 | UV.UV_VIS_2.AcqOn
12:29:02 | 0.000 | Entered stage "Run"
12:29:32 | 0.500 | Log Background: 0.2 [µS]
12:29:32 | 0.500 | Log Pressure: 2999.8 [psi]
12:34:02 | 5.000 | DC.Compartment_TC.ReleaseExclusiveAccess
12:34:02 | 5.000 | Exclusive access finished.
12:34:02 | 5.000 | DC.Column_TC.ReleaseExclusiveAccess
12:34:02 | 5.000 | Exclusive access finished.
12:34:02 | 5.000 | Exclusive access finished.
12:44:02 | 15.000 | Entered stage "Stop Run"
12:44:02 | 15.000 | UV.UV_VIS_1.AcqOff
12:44:02 | 15.000 | CDet1.CD_1.AcqOff
12:44:02 | 15.000 | CDet1.CD_1_Total.AcqOff
12:44:02 | 15.000 | UV.UV_VIS_2.AcqOff
12:44:02 | 15.000 | Entered stage "Post Run"
12:44:02 | 15.000 | DC.Column_TC.ReleaseExclusiveAccess
12:44:02 | 15.000 | DC.Compartment_TC.ReleaseExclusiveAccess
12:44:02 | 15.000 | Terminating instrument method chrom://desktop-5bv84qa/ChromeleonLocal/Instrument Data/SYSTEM_1/60511 Compass/091522.seq/AS11HC_SM_UV.instmeth. Waiting for instrument to finish processing of current injection ...
12:44:03 | 15.016 | OK to terminate instrument method chrom://desktop-5bv84qa/ChromeleonLocal/Instrument Data/SYSTEM_1/60511 Compass/091522.seq/AS11HC_SM_UV.instmeth.
12:44:04 |  | End of injection "1".